## Slide 1: iTPM
Monica Alderighi, monica@iasf-milano.inaf.it
Sandro Pastore, sandro.pastore@sanitaseg.it
SKA AADC All Hands Meeting, May 9th - 13th, 2016

## Slide 2: Who we are
Team @INAF-IASF Milano
  2 Staff, 5 Research Associates, 4 Collaborators, Bachelor/Master students
  High reliability computing systems for Astrophysical apps
Sanitas EG s.r.l.
  Founded by former students
  Historical partner of IASF Milano
    Scientific collaboration
    Prototypes
  Industrial products
    Single board computers
    Cameras
    FPGA & SW IP CORE

## Slide 3: Beginning…
INAF project approved for funding, May 2013
“Digital Platform development for back end design of new generation SKA Aperture Arrays”
  2 year program
  P.I. Francesco Schillirò (OACT)
  5 INAF Institutes involved
Accepted within LFAA AADC, December 2013 (Signal Processing WP)
INAF- Arcetri
INAF IASF
Milano
INAF IRA
Medicina
INAF- OA
Cagliari
INAF –OA
Catania

## Slide 4: AADC Participants and Partners

## Slide 5: Initial board proposal
Handling of 8-12-16 antennas
Analogue signal acquisition through 8 bit ADC @700 MHz
Two polarization signals
In calibration it must be possible to transfer the whole acq BW
Input signals must be filtered to extract 1Mhz frequency bands
  1 Mbit for each channels
Giga Ethernet protocol for the output communication (likely 40G via QSFP or similar)
MODULAR
&
SCALABLE

## Slide 6: iTPM in LFAA
Tile: 16 Antennas, 2 Pol, 32 Channels
AADC-TEL.LFAA.SE.MGT-AADC-PL-002
Low Frequency Aperture Array Technical Description, Issue 1, 2013
Tile processor implementation

## Slide 7: iTPM-ADU first concept, Jan 2014
Main function co-existence:
  Online digital data processing: 2 large FPGAs
  Analog to digital low noise conversion
    32 analog input
    Up to 1 Giga sample at 8 bit
  Board management
Mechanical specs: max 4U
Total power < 90 Watt
Cost, testability & manufacturability
Feasibility study

## Slide 8: Feasibility study topics
Main device: FPGA & AD
Rack assembly
Board size & Power dissipation
Board technology
  PCB build-up
  Isolation criteria
  Sampling clock architecture

## Slide 9: Feasibility output: main devices
Direct support from market leader suppliers
  Xilinx: state of the art 20nm Ultrascale FPGA, and early access program
  Analog Device: last devices close to release
    JESDB 14 bit 1 GHz dual channel AD
    JESDB PLL
    Ultra High PSRR linear regulator for AD
    SKA DEMO FMC board to performance evaluation
Notes: Power Supply Rejection Ratio

## Slide 10: FPGA Xilinx Ultrascale: 20 nm device
Production 1q 2015
Block RAM Capacity (Mb):
DSP Slice Count
Serial Transceivers:

## Slide 11: AD9680 Datasheet
Input 0
Input 1
Clock
JESD IF
700 mW
per channel
1q 2016
1.6 W
per channel
??? mW
per channel

## Slide 12: Feasibility output: preliminary board design
  Board architecture
  Expected performances
  Expected power consumption
  Costs & production proc
  Production time line

## Slide 13: Performed activities
Feasibility Study: January 2014
Prototype production
  R 1.0 (#2): January 2015
  R 1.1 (#4): October 2015
Performance Tests
  Data acquisition, configurations: Spring 2015
  Power consumption and temperature in rack: December 2015

## Slide 14: Current & future activities
iTPM-ADU Public Tender for AAVS1
  Published April 18, 2016
  Closure May 18, 2016
Testing & AAVS1 production support
  Test Specification
  Test Firmware
  Procedures for functional verification of the iTPM ADU board
  Test JIG development
iTPM-ADU development
  R 2.0 prototype featuring low power ADCs
  High density rack capability

## Slide 15: Activity Breakdown
15
iTPM Board Development
Management
HW + Basic Firmware
Performance Test
Operation Firmware
Test Bench
Mechanical Design
Software
OACT/IASF
IASF/Sanitas EG
Arcetri/Sanitas EG/ Oxford/RAL
IASF/Sanitas EG
IRA
IRA/Sanitas EG
Malta Univ/Oxford

## Slide 16: iTPM key

## Slide 17
DDR
DDR
DDR
DDR
XO2
CPU
QSFP+
QSFP+
RJ45
AD
AD
AD
AD
AD
AD
AD
AD
AD
AD
AD
AD
AD
AD
AD
AD
PLL
ADA
ADA
ADA
ADA
ADA
ADA
IJ5
IJ5
IJ5
IJ5
ADA
ADA
ADA
ADA
ADA
ADA
ADA
ADA
CLKB
SYSB
CLKB
SYSB
LAY 2-11
100 Ω
ADD
ADD
INT BUS
Board
Placement issues
Input DC-DC
DC-DC 4-13
DC-DC 1-3
DC-DC 14-17
DIGITAL SIDE

## Slide 18: ADC related issues (partial)
Device specific positioning and layout issues:
  HS signal routing (16 device in 233 mm board height)
  Staggered placement
  Clock and Sync routing
Output JESD
AD9680
Power and
decoupling
Input RF
Syncro signals
8 GND & Supply
Analog Inputs
12.5 Gbps
digital out
AD9680
AD9680
AD9680
Clock & syncro

## Slide 19: Stripline technology, 10 G JESDB, 10Geth, Analog Signal, Analog clocks
Stripline on Layer 2, 13
  External shield, layer 1, 14
  Reference GND plane, layer 3, 12
Reference Ground
Signal
path

## Slide 20: Analogue input & JESD clock layout

## Slide 21: iTPM features

## Slide 22: iTPM board
Compact size, 6U format compatible
32 analog input capability with high quality and high isolation connectors
32 analog programmable driver (amplifier/attenuator)
16 dual channel 14 bit 1Ghz AD converter
High quality clock distribution system JESD204B Subclass 1 with deterministic latency
Two Xilinx Ultrascale XCU40 20 nm FPGAs dedicated to data processing
Two DDR3, 96 bit memory banks, 6+6 Gbit total size, 1800 Mhz speed
Two 40G Ethernet interface (QSFP)
FPGAs high speed internal connection bus, 25 Gbit + 25 Gbit bidirectional
Ethernet gigabit management with instant ON flash based CPLD
Complete system monitor interface with power, voltages and temperatures monitoring
On board user flash for multiple FPGA configurations options
Management microprocessor (ARM Cortex A8, 32 bit, 240 Mhz)
High efficient multistage configurable Power Distribution system

## Slide 23: Acquisition and processing architecture
FPGA USAGE: | JESD RX | INT BUS | DDR | 40G Eth
 | 16 GByte/s | 5.0 GByte/s | 16 GByte/s | 5.0 GByte/s

## Slide 24: Mechanical specification
6U high and 170 mm depth
standard rack mount format
Top mount heat sink

## Slide 25: Management
  Supply control
  Voltage, current and temperature monitoring
  Ethernet GIGA management interface
  Board configuration
    Configuration signals and device register
    FLASH Memory read/write
    FPGA configuration
    Firmware Update (multiple bitstream)
    FPGA scrubbing (SEE monitoring)
  CPU complex procedure execution
    TCP protocol, 100 Mbit
    Complex management procedure
DEBUG & TEST
INTEGRATION

## Slide 26: Monitoring
Voltages: 40 point of measure:
  Input
  All DC-DC
  AD linear filtered group
Current: 3 point of measure:
  Input
  FE output
Temperature: 4 point of measure:
  Board management
  Board analog area
  FPGAs core

## Slide 27: Data synchronization JESD204B, sub1
Ext source
clock
Clock network
On board
Oscillator
RX JesdB
RX JesdB
RX JesdB
SYSREF Gen
Dev clock
Dev clock side 0
FPGA 0
AD side 0
SIDE 1
SIDE 0
SYSREF side 0
FPGA 1
PPS
Internal
PLL
Notes: SYSREF must be source synchronous with the device clock

## Slide 28: Power consumption

## Slide 29: iTPM Numbers
Placement Areas: Available: 123,45 Sq. (in)
  Required: 37,99 Sq. (in)
  Required/Available: 30.78 %
Connections ...................... 13.312
Differential Pairs ............ 413
Total Trace Length ............ 233 meter
Vias ............................... 10.800
Parts Placed ..................... 4.214
  Parts Mounted on Top … 2.159
  Parts Mounted on Bottom … 2.055

## Slide 30: RF qualification
TPM TEST procedure

## Slide 31: RF qualification
Test objective: verify iTPM RF signal acquisition capability and quality.
Band selection, iTPM-ADU sampling clock:
  50-375, sampling clock 800 MHz
  375-650, sampling clock 700 MHz
  TBD 1000 MHz
Single channel RAW acquisition for selected input frequencies (50-60 frequency point)
Multiple channel RAW acquisition for cross talk measurement (TBD frequency point)
S11 measure ??? - OFFLINE -

## Slide 32: ADU Board Test Setup
ADU Board Test Setup
Notes: In the picture there are the elements of the system:
A Signal Generator, with low harmonics, good spectral purity, low phase noise
Filters to suppress harmonics and spurious signals
Power meter to measure the RF output power of Signal Generator and for its control
Source and Sinks must be locked to the same stable Reference
A workstation that controls the instruments over a GPIB bus, controls also the ADU board, receives and processes measurements data

## Slide 33: Spurious Free Dynamic Range
375-650 MHz
50-375 MHz
Notes: I show you some results from the measurement on the BOARD with ADA.
The dispersion of all the input trends is within 0.6 dB. Also the gain flatness is within 0.6 dB

## Slide 34: Effective Number of Bit
50-375 MHz
375-650 MHz
Notes: I show you some results from the measurement on the BOARD with ADA.
The dispersion of all the input trends is within 0.6 dB. Also the gain flatness is within 0.6 dB

## Slide 35: Cross-Talk Measurement
#0
#1
#3
#2
#5
#7
#9
#11
#13
#15
#4
#6
#8
#10
#12
#14
Input numbering

## Slide 36: Cross-Talk Measurement
BANDWIDTH: 50-375 MHz, signal injected in Input 4, no heat sink shielding
WORST CASE

## Slide 37: Return Loss
50-375 MHz
375-650 MHz

## Slide 38: Measurement Summary
NOTE: all values are based on WORST CASE
Fs: 800 MSPS --- BW: 50 – 375 MHz |  | 
ADC Performance Parameters | ADU Board#1 (with ADA) | ADU Board#2 (without ADA)
Signal to Noise Ratio referenced to Full Scale [dBFS] | ≥ 49.19 | ≥ 49.33
Gain Flatness [dBFS] | ≤ ±0.3573 | ≤ ±0.343
2nd-order Harmonic Distortion [dBc] | ≤ -67.24 | ≤ -67.74
3rd-order Harmonic Distortion [dBc] | ≤ -66.53 | ≤ -68.56
Worst Other Spur [dBc] | ≤ -67.03 | ≤ -66.83
Spurious Free Dynamic Range [dBc] | ≥ 66.53 | ≥ 66.83
ENOB [bits] | ≥ 7.876 | ≥ 7.896
Cross-Talk [dBc] | ≤ -65.69 | ≤ -61
IP3 [dB] (F1=184.7 MHz; F2=187.5 MHz) | 29.55 | 32.2
IP2 [dB] (F1=184.7 MHz; F2=187.5 MHz) | 66.3 | 77.5
Fs: 700 MSPS --- BW: 375 – 650 MHz |  | 
ADC Performance Parameters | ADU Board#1 (with ADA) | ADU Board#2 (without ADA)
Signal to Noise Ratio referenced to Full Scale [dBFS] | ≥ 48.88 | ≥ 49.32
Gain Flatness [dBFS] | ≤ ±0.6252 | ≤ ±1.356
2nd-order Harmonic Distortion [dBc] | ≤ -65.77 | ≤ -59.9
3rd-order Harmonic Distortion [dBc] | ≤ -60.78 | ≤ -65.64
Worst Other Spur [dBc] | ≤ -64.16 | ≤ -63.23
Spurious Free Dynamic Range [dBc] | ≥ 60.78 | ≥ 59.9
ENOB [bits] | ≥ 7.788 | ≥ 7.886
Cross-Talk [dBc] | ≤ -70.58 | ≤ -70.39
IP3 [dB] (F1=500.1 MHz; F2=503.2 MHz) | 26 | 24.4
IP2 [dB] (F1=500.1 MHz; F2=503.2 MHz) | 64 | 58.3

## Slide 39: iTPM qualification & testing

## Slide 40: Test Objectives
Production test
  Fast and simple
  Identify main electrical production defects
Functional test
  Single wire oriented to guide defect identification
  Board delivery
    Complete functional verification
    High speed qualification
    Acceptance procedure report generation
Qualification test
  SKA installation
    RF test
    Integration test

## Slide 41: iTPM Test strategy
Three level testing
  ON line production testing
  Complete functional test
    Management test
    Digital Test
    Functional analog test
  Quality analog test
    RF full qualification
Different place, skills & instrument needed
SKA delivery
Production delivery

## Slide 42: iTPM Test Implementation
Production test plant
QUALIFICATION TEST PLANT
FUNCTIONAL TEST PLANT
iTPM TEST JIG
iTPM RF JIG
“SKA quality” RF signal generators
“Medium quality” RF signal generators
SERVER data analysis
iTPM TEST JIG control SW & CPLD-CPU FW
iTPM FPGA TEST FW
iTPM FPGA SKA FW
iTPM Production SW
TPM Qualif. SW
HW
FW & SW

## Slide 43: Test requirement
Complete functional check, digital devices
  Power check
Multi gigabit data lines check
  Electrical margins
  BER check
Full load board
  Power Noise impact on board function
  Noise vs. board temperature
Analog interface
  Functional & RF qualification

## Slide 44: Functional & Qualification set up
32
16
40 G
16
40 G
RF
JIG
Signal
generator
40 G
40 G
1 G
1 G
1 G
RAM
X 4
CPU
DISC
USER INTERFACE (TBC)
MEASURE SET UP
RAW DATA RAM BUFFERING
DATA ANALISYS
QUALIFICATION MANAGER
QUALIFICATION
FW FUNCTION
SERVER
iTPM
JIG
POWER
SUPPLY
HW
SW - FW

## Slide 45: RF JIG Architecture
RF FILTERS
7x
ADU RF INPUT
ADU RF INPUT
CPLD
Eth UDP
Power Meter
PLL
LDO
Gain
1
PPS
SPI IF
LOAD
8 Watt
ADU IF
10 Mhz
FAN IF
2
32 channel path

## Slide 46: Single signal path iTPM RF generator:
iTPM RF INPUT
iTPM RF INPUT
RF FILTERS
7x
-1 dB FS
Power Meter
Coupler
Gain
PLL
2
-1.5 dB
-2.0 dB
-3.5 dB
-6,0dB (worst)
+11.8dB
+15.6 .. +18.1dB
Att
Fast switching time
High power
iP3

## Slide 47: Filter selection
All COTS filters (except KR228975S)
Parameters considered :
Freq Band: 50÷1250 MHz
Rejection at 2f ≥ 40dB
Rejection at f/2 ≥ 20dB
In band IL ≤ 2.5dB
In band Return Loss ≥ 15dB

## Slide 48: RF performances
Isolation between channels > 80dB
IRL&ORL for active channel better than 10dB
ORL for non-active channels ≥ 20dB

## Slide 49: iTPM Next release: 2.0
New AD
  Pin compatible, same performance
  Sample rate up to 1.250 Msample/sec
  Less power, more than 50%
New Amplifier, half power
New PLL, Jitter reduction
DDR4 Memory
  Power reduction, 40% expected
  Double capacity, Speed increase
Backplane interface ?

## Slide 50: Thanks !!!

## Slide 51: RF Qualification FW function
All single channel measures:
  RAW data capture, selectable depth single channel, 8 bit mode
    1 MB (Required)
    2 MB (TBC)
    More MB ? (TBD)
  RAW data capture, 12/14 bit mode (TBD)
Multichannel measures:
  RAW data capture, 2 MB total buffer for 16 input, selectable n° of adjacent input
    16 input (ALL), 128 KB RAW each
    8 input, 512 KB RAW each
Depth: | Frequency resolution | 
RAW KB | Low BW, 50-375 Mhz | High BW, 375-650 Mhz
128 | 6.10 KHz | 5.34 Khz
512 | 1.56 KHz | 1.33 Khz
1024 | 762.9 Hz | 667.6 Hz
2048 | 381.5 Hz | 333.8 Hz
16384 | 47,6 Hz | 41,7 Hz
Current spectra resolution of the CSP: 1 KHz

## Slide 52: RF qualification: measures calculation
RAW buffering
10 x 1 Mega samples
1M samples
1M samples
1M samples
FFT
FFT
FFT
MEDIA
Matlab Operators
Measure
SET UP
And
X10
acquisition
Preliminary evaluation
Data analysis
Compliance check
Measures

## Slide 53: ITPM Test Implementation
Hardware
  Test JIG (Production and Functional)
  Ad hoc JIG (Qualification)
    RF signal generator
  Server for data analysis
Firmware
  FPGA test firmware (Production, Functional, Qualification)
Software
  Production, Functional, Qualification

## Slide 54: Xilinx test FW option 2: SDK integration
Management
BUS
JESD IF
Link 0-7
JESD IF
Link 8-15
4x10G Eth
Link 16-19
Data
buffer
FFT
Select I/O
16 + HSQ
Select I/O
16 + HSQ
Data
buffer
FFT
DRP
DRP
DRP
DDR IF
XSDB
Microblaze
32 Bit
200 Mhz
AXI DMA
AXI BUS TEST
AXI BUS OXFORD
System monitor
Master
Slave
Slave
Slave
Master
BRam